## Slide 1: Paul Scherrer Institute, Switzerland
Fast Wave-form Sampling Front-end Electronics
Stefan Ritt

## Slide 2: Prologue
RTSD Luncheon

## Slide 5: Undersampling of signals
Undersampling: Acquisition of signals with sampling rates ≪ 2 * highest frequency in signal
Image Processing
Waveform Processing

## Slide 6: Agenda
1
What is the problem?
2
Tool to solve it
3
What else can we
do with that tool?

## Slide 7
1
What is the problem?

## Slide 8: Signals in particle physics
Photomultiplier (PMT)
Scintillator
Particle
10 – 100 ns
HV
HV
1 – 10 ms
Scintillators
(Plastic, Crystals, Noble Liquids, …)
Wire chambers Straw tubes
Silicon
Germanium

## Slide 9: Measure precise timing: ToF-PET
Positron Emission Tomography
Time-of-Flight PET
Dt
d
d ~ c/2 * Dt
e.g. d=1 cm → Dt = 67 ps

## Slide 10: Traditional DAQ in Particle Physics
Threshold
Threshold
TDC
(Clock)
+
-
ADC
~MHz

## Slide 11: Signal discrimination
Threshold
Single Threshold
“Time-Walk”
Multiple Thresholds
T1
T2
T3
T1
T2
T3
Inverter & Attenuator
S
Delay
Adder
0
Constant Fraction (CFD)

## Slide 12: Influence of noise
Voltage noise causes
timing jitter !
Fourier Spectrum
Signal
Noise
Low pass filter
Low pass filter (shaper) reduces noise while maintaining most of the signal

## Slide 13: Noise limited time accuracy
U [mV] | DU [mV] | tr | Dt
100 | 1 | 1 ns | 10 ps
10 | 1 | 3 ns | 300 ps
All values in this talk are s (RMS) !
FHWM = 2.35 x s
Most today’s TDCs have ~20 ps LSB
How can we do better ?

## Slide 14: Noise limited time accuracy

## Slide 15: Switching to Waveform digitizing
ADC
~100 MHz
FPGA
Advantages:
General trend in signal processing (“Software Defined Radio”)
Less hardware (Only ADC and FPGA)
Algorithms can be complex (peak finding, peak counting, waveform fitting)
Algorithms can be changed without changing the hardware
Storage of full waveforms allow elaborate offline analysis
SDR

## Slide 16: Example: CFG in FPGA
+
Adder
Look-up Table
(LUT)
8-bit address
8-bit data
* (-0.3)
S
Delay
Adder
0
Latch
Clock
>0
≤ 0
AND
Delay
FPGA

## Slide 17: Nyquist-Shannon Sampling Theorem
fsignal < fsampling /2
fsignal > fsampling /2

## Slide 18: Limits of waveform digitizing
Aliasing Occurs if fsignal > 0.5 * fsampling
Features of the signal can be lost (“pile-up”)
Measurement of time becomes hard
ADC resolution limits energy measurement
Need very fast high resolution ADC

## Slide 19: What are the fastest detectors?
Micro-Channel-Plates (MCP)
      Photomultipliers with thousands of tiny channels (3-10 mm)
      Typical gain of 10,000 per plate
      Very fast rise time down to 70 ps
70 ps rise time  4-5 GHz BW  10 GSPS
SiPMs (Silicon PMTs) are also getting < 100 ps
J. Milnes, J. Howoth, Photek

## Slide 20: Can it be done with FADCs?
8 bits – 3 GS/s – 1.9 W  24 Gbits/s
10 bits – 3 GS/s – 3.6 W  30 Gbits/s
12 bits – 3.6 GS/s – 3.9 W  43.2 Gbits/s
14 bits – 0.4 GS/s – 2.5 W  5.6 Gbits/s
1.8 GHz!
24x1.8 Gbits/s
Requires high-end FPGA
Complex board design
High FPGA power
PX1500-4: 2 Channel 3 GS/s 8 bits
ADC12D1X00RB: 1 Channel 1.8 GS/s 12 bits
1-10 k$ / channel
What about 1000+ Channels?
V1761: 2 Channels, 4 GS/s, 10 bits

## Slide 21
2
Tool to solve it

## Slide 22: Switched Capacitor Array (Analog Memory)
Shift Register
Clock
IN
Out
“Time stretcher” GHz  MHz
Waveform
stored
Inverter “Domino” ring chain
0.2-2 ns
FADC
33 MHz
10-100 mW

## Slide 23: Time Stretch Ratio (TSR)
Clock
IN
Out
dts
dtd
Typical values:
dts = 0.5 ns (2 GSPS)
dtd = 30 ns (33 MHz) → TSR = 60

## Slide 24: Triggered Operation
sampling
digitization
lost events
sampling
digitization
Sampling Windows * TSR
Chips usually cannot sample during readout ⇒ “Dead Time”
Technique only works for “events” and “triggers”
Dead time =
Sampling Window ∙ TSR
(e.g. 100 ns ∙ 60 = 6 ms)

## Slide 25: Time resolution limit of SCA
PCB
Det.
Chip
Cpar

## Slide 26: Bandwidth STURM2 (32 sampling cells)
G. Varner, Dec. 2009

## Slide 27: How is timing resolution affected?
U | Du | fs | f3db | Dt
100 mV | 1 mV | 2 GSPS | 300 MHz | ∼10 ps
1 V | 1 mV | 2 GSPS | 300 MHz | 1 ps
100 mV | 1 mV | 10 GSPS | 3 GHz | 1 ps
today:
optimized SNR:
next generation:
- high frequency noise
- quantization noise

## Slide 28: Timing Nonlinearity
Bin-to-bin variation: “differential timing nonlinearity”
Difference along the whole chip: “integral timing nonlinearity”
Nonlinearity comes from size (doping) of inverters and is stable over time → can be calibrated
Residual random jitter: 1-2 ps RMS beats best TDC
Recently achieved with new calibration method http://arxiv.org/abs/1405.4975
Dt
Dt
Dt
Dt
Dt

## Slide 29: First Switched Capacitor Arrays
IEEE Transactions on Nuclear Science,
Vol. 35, No. 1, Feb. 1988
50 MSPS in 3.5 mm CMOS process

## Slide 30: Switched Capacitor Arrays for Particle Physics
STRAW3
TARGET
LABRADOR3
AFTER
NECTAR0
SAM
E. Delagnes
D. Breton
CEA Saclay
DRS1
DRS2
DRS3
DRS4
G. Varner, Univ. of Hawaii
0.25 mm TSMC
Many chips for different projects (Belle, Anita, IceCube …)
0.35 mm AMS
T2K TPC, Antares, Hess2, CTA
H. Frisch et al., Univ. Chicago
PSEC1 - PSEC4
0.13 mm IBM
Large Area Picosecond Photo-Detectors Project (LAPPD)
2002
2004
2007
2008
0.25 mm UMC
Universal chip for many applications
MEG experiment, MAGIC, Veritas, TOF-PET
SR
R. Dinapoli
PSI, Switzerland
drs.web.psi.ch
www.phys.hawaii.edu/~idlab/
matacq.free.fr
psec.uchicago.edu

## Slide 31: Some specialities
LAB Chip Family (G. Varner)
  Deep buffer (BLAB Chip: 64k)
  Double buffer readout (LAB4)
  Wilkinson ADC
NECTAR0 Chip (E. Delagnes)
  Matrix layout (short inverter chain)
  Input buffer (300-400 MHz)
  Large storage cell (>12 bit SNR)
  20 MHz pipeline ADC on chip
PSEC4 Chip (E. Oberla, H. Grabas)
  15 GSPS
  1.6 GHz BW @ 256 cells
  Wilkinson ADC
6 mm
16 mm
Wilkinson-ADC:
Ramp
Cell contents
measure time

## Slide 32
3
What can we do with that tool?

## Slide 33: MEG On-line waveform display
template
fit
S848
PMTs
“virtual oscilloscope”
Liq. Xe
PMT
1.5m
g
m+e+g
At 10-13 level
3000 Channels
Digitized with DRS4 chips at
1.6 GSPS
m
Drawback: 400 TB data/year

## Slide 34: Pulse shape discrimination
g
a
m
g
a
m
Events found and correctly processed 2 years (!) after the were acquired

## Slide 35: Readout of Straw Tubes
HV
Readout of straw tubes or drift chambers usually with “charge sharing”: 1-2 cm resolution
Readout with fast timing: 10 ps / √10 = 3 ps → 0.5 mm
Currently ongoing research project at PSI
d ~ c/2 * Dt

## Slide 36: A first test
Speed: 266 mm/ns (7.5 ps/mm)
Accuracy: 4.2 ps or 0.5 mm

## Slide 37: MAGIC Telescope
http://ihp-lx.ethz.ch/Stamet/magic/magicIntro.html
https://wwwmagic.mpp.mpg.de/
La Palma, Canary Islands, Spain, 2200 m above sea level

## Slide 38: MAGIC Readout Electronics
Old system:
2 GHz flash (multiplexed)
512 channels
Total of five racks, ~20 kW
New system:
2 GHz SCA (DRS4 based)
2000 channels
4 VME crates
Channel density 10x higher

## Slide 39: Digital Pulse Processing (DPP)
C. Tintori (CAEN)
V. Jordanov et al., NIM A353, 261 (1994)

## Slide 40: Template Fit
Determine “standard” PMT pulse by averaging over many events  “Template”
  Find hit in waveform
  Shift (“TDC”) and scale (“ADC”) template to hit
  Minimize c2
  Compare fit with waveform
  Repeat if above threshold
Store ADC & TDC values
pb Experiment
500 MHz sampling
www.southerninnovation.com
14 bit
60 MHz

## Slide 41: Other Applications
Gamma-ray astronomy
Magic
CTA
Antarctic Impulsive Transient Antenna
(ANITA)
320 ps
IceCube
(Antarctica)
Antares
(Mediterranian)
ToF PET (Siemens)

## Slide 42: High speed USB oscilloscope
4 channels
5 GSPS
1 GHz BW
8 bit (6-7)
15k€
4 channels
5 GSPS
1 GHz BW
11.5 bits
900€
USB Power
Demo

## Slide 43: SCA Usage

## Slide 44: Things you can buy
DRS4 chip (PSI)
32+2 channels
12 bit 5 GSPS
> 500 MHz analog BW
1024 sample points/chn.
110 ms dead time
MATACQ chip (CEA/IN2P3)
4 channels
14 bit 2 GSPS
300 MHz analog BW
2520 sample points/chn.
650 ms dead time
DRS4 Evaluation Board
4 channels
12 bit 5 GSPS
750 MHz analog BW
1024 sample points/chn.
500 events/sec over USB 2.0
SAM Chip (CEA/IN2PD)
2 channels
12 bit 3.2 GSPS
300 MHz analog BW
256 sample points/chn.
On-board spectroscopy

## Slide 45: Next Generation SCA
Low parasitic input capacitance
Wide input bus
Low Ron write switches  High bandwidth
Short sampling depth
Digitize long waveforms
Accommodate long trigger delay
Faster sampling speed for a given trigger latency
Deep sampling depth
How to combine
best of both worlds?

## Slide 46: Cascaded Switched Capacitor Arrays
shift register
input
fast sampling stage
secondary sampling stage
. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .
32 fast sampling cells (10 GSPS)
100 ps sample time, 3.1 ns hold time
Hold time long enough to transfer voltage to secondary sampling stage with moderately fast buffer (300 MHz)
Shift register gets clocked by inverter chain from fast sampling stage

## Slide 47: The dead-time problem
Only short segments of waveform are of interest
sampling
digitization
lost events
sampling
digitization
Sampling Windows * TSR

## Slide 48: FIFO-type analog sampler
digitization
FIFO sampler becomes immediately active after hit
Samples are digitized asynchronously
“De-randomization” of data
Can work dead-time less up to average rate = 1/(window size * TSR)
Example: 2 GSPS, 10 ns window size, TSR = 60 → rate up to 1.6 MHz

## Slide 49: DRS5
Self-trigger writing of 128 short 32-bin segments (4096 bins total)
Storage of 128 events
  Accommodate long trigger latencies
  Quasi dead time-free up to a few MHz,
  Possibility to skip segments → second level trigger
Attractive replacement for CFG+TDC
Delay chain tested in 0.11 mm UMC process
First version planned for 2016

## Slide 50: SAMPIC Chip (E. Delagnes et al)
“Waveform TDC”: Coarse timing by TDC + interpolation by waveform digitizing of 64 analog sampling cells + ADC readout
Simultaneous write & read
5 ps (RMS) time resolution at 2 MHz event rate
Planned for ATLAS AFP and SuperB TOF

## Slide 51: Conclusions
SCA technology offers tremendous opportunities
Several chips and boards are on the market for evaluation
New series of chips on the horizon might change front-end electronics significantly